Authorized by:

Jaime Nadal

UNFPA Representative

Date: 29/06/2023

Request for Quotation No. UNFPA/UKR/RFQ/23/21

Dear Sir/Madam,

We hereby solicit your quotation for the supply of the items named below and their LAND shipment to Kyiv, Ukraine (CPT Kyiv):

| Lot No. and Item | Product Description | Quantity |
| --- | --- | --- |
| 1. Colposcope | • Magnification: A range of optical magnification between 3x to 15x (either stepped or continuously variable) • Illumination: Light source LED 20,000-35,000 LUX (at 300 mm working distance) - Zoom, green filter, focus manually and adjust light brightness by hand controller • Ingress protection rating: IPX2 (minimum) Software Colposcope software should be in English as mandatory and Russian (Ukrainian language, if available) - Data storage: 200 or more - standard pictures - Multi-model (case number, name, age, date, etc) index, lookup, analyze, edit and print. Printing of Multi-format diagnosis report - Acetic acid test timer function, real-time display of acetic iodine reaction tests. Parameters: Camera System：1/4 inch. Fast auto focus/manual focus/ Effective Pixel：>=800,000 pixels Resolution: 550 lines Signal Noise Ratio of lens: >=50dB Focusing distance: 20-40mm Electrical Requirements Voltage: 230 V, 50 Hz Plug type: C or F All electrical connection according with the destination Country standards (Ukraine) Accessories: - Upright Stand, easy to move, with fastening to a chair to allow for hands-free operation • LED TV or medical grade monitor if not integrated (optional) Documentation requirement: • Instructions for use and service manuals to be provided (including procedures for decontamination) • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local service agent. • Certificate of calibration and inspection to be provided • List to be provided of equipment and procedures required for local calibration and routine maintenance. • List to be provided of common spares and accessories, with part numbers. Other requirements. All standard accessories, consumables and parts required to operate the equipment, including all standard tools, cleaning and lubrication materials. At least two (2) years of full onsite warranty including 2 yearly preventive maintenances, full corrective maintenance including works and spare parts. Training on use, cleaning, disinfecting and basic maintenance on-site or in teleconference for at least 4 hours. Regulatory approvals: Compliance to - National Regulatory Agency/Authority (NRA) requirements compliance And at least one of: - FDA 510k clearance (US FDA) - CE mark (EU), with indication of Notifying Body - Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan. Manufacturer authorization QMS For general manufacturing • ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes • ISO 14971: Medical Devices - Application of Risk Management to Medical Devices • ISO 15223-1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements Safety and product standards: Declaration of Conformity by the Manufacturer is acceptable • IEC 60601-1 - Medical electrical equipment - Part 1: General requirements for basic safety and essential performance • IEC 60601-1-2: Medical electrical equipment - Part 1-2 General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests For vaginally-inserted colposcopes: • Biocompatibility: Declaration of Conformity by the Manufacturer is acceptable o ISO 10993-1: Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process o ISO 10993-5: Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity o ISO 10993-10: Biological evaluation of medical devices Part 10: Tests for irritation and skin sensitization • Endoscopy (pertaining to inserted scope): o ISO 8600-1: Endoscopes – Medical endoscopes and endotherapy devices – Part 1: General requirements o ISO 8600-3: Optics and optical instruments – Medical endoscopes and endoscopic accessories – Part 3: Determination of field of view and direction of view of endoscopes with optics o ISO 8600-4: Endoscopes – Medical endoscopes and endotherapy devices – Part 4: Determination of maximum width of insertion portion o ISO 8600-5: Optics and photonics – Medical endoscopes and endotherapy devices – Part 5: Determination of optical resolution of rigid endoscopes with optics o ISO 8600-6: Optics and photonics – Medical endoscopes and endotherapy devices – Part 6: vocabulary | 7 pcs |
| 2. 12 channel ECG device | 12-channels ECG portable device used for Electrocardiogram (ECG) examinations. General Specifications: 12 electrographic standard leads measuring with simultaneous acquisition. Visualization of at least one group of 3 leads simultaneously. At least one ECG patient cable with at least 5 terminals. At least 8” LCD or TFT color display with visualization of: analogical curves, alphanumeric values and the related physiological limits. Screen resolution >=128 х 64 pixels; Keyboard- combined alphanumeric and functional keyboard and buttons; Number of leads 12; Bandwidth in monitor mode not smaller than 0.5 to 100 Hz Modes of operation: automatic and manual. Equipment compatible with patients with pacemakers. All software in English (as mandatory) and Russian, Ukrainian language (if available). Protection against defibrillation. Automatic internal data storing up to at least 40 ECG records. Input impedance not less than 5MΩ. Patient leakage current no higher than 10uA. Common Mode Rejection (CMRR) not less than +100dB. ECG signal measurement range not smaller than -2 mV to +2 mV. At least the following adjustable alarms: Heart failure; Ventricular fibrillation; Tachycardia; Bradycardia; Electrode disconnection; Open lead; Low battery visual alarm; Equipment provided with at least the following software applications: Arrhythmias detection and analysis; Ventricular fibrillation detection and analysis; Ventricular tachycardia detection and analysis; ST segment analysis; Automatic equipment calibration. RS232 or USB port for PC connection and data transmission, WiFi (optional). With filters at least for baseline instability filter muscle artifacts, AC interference and other devices. Printer Integrated printer using standard A4 format paper and different report print formats available. Print type - Thermal Paper type - roll Selectable printing paper speed at least of 5, 25 and 50 mm/sec. Electrical requirements: 210 to 240 V, 50 Hz Plug type: F Integrated batteries charger with AC power cable and rechargeable batteries with continuous monitoring working time of no less than 1 hour or 100 ECG. Protections against over-voltage and over-current line conditions. All electrical connection according with the destination Country standards (Ukraine) Accessories to be included: Two (2) lead patient cables with at least five terminals, reusable type Eight (8) reusable clip clamp electrodes, reusable type Two sets of twelve (12) reusable suction ball-type chest electrodes, reusable type Two (2) bottles of gel for body-connection electrodes. Two (2) rolls of ECG thermal paper. Documentation requirements • Software licenses and user and maintenance manuals. Instructions for use and service manuals to be provided (including procedures for decontamination) • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local service agent. • Certificate of calibration and inspection to be provided • List to be provided of equipment and procedures required for local calibration and routine maintenance. • List to be provided of common spare and accessories, with part numbers. Other requirements. All standard accessories, consumables and parts required to operate the equipment, including all standard tools, cleaning and lubrication materials. At least two (2) years of full onsite warranty including 2 yearly preventive maintenances, full corrective maintenance including works and spare parts. Training on use, cleaning, disinfecting and basic maintenance on-site or in teleconference for at least 4 hours. Regulatory approvals: Compliance to - National Regulatory Agency/Authority (NRA) requirements compliance Free Sales certificate And at least one of: - FDA 510k clearance (US FDA) - CE mark (EU), with indication of Notifying Body - Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan. Manufacturer authorization QMS ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes). General quality management (e.g. ISO 9001:2015 Quality management systems – Requirements). Application of risk management to medical devices (e.g. ISO 14971:2019 Medical devices – Application of risk management to medical devices). Safety and product standards: Declaration of Conformity by the Manufacturer is acceptable IEC 60601-1:2012 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-2:2007 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. IEC 60601-1-8:2012 Medical electrical equipment – Part 1-8: General requirements for basic safety and essential performance – Collateral standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems. IEC 60601-2-25:2011 Part 2-25: Particular requirements for the basic safety and essential performance of electrocardiographs. IEC 60601-2-27:2011 Part 2-27: Particular requirements for the basic safety and essential performance of electrocardiographic monitoring equipment. IEC 60601-2-49:2011 Part 2-49: Particular requirements for the basic safety and essential performance of multifunction patient monitoring equipment. | 7 pcs |
| 3. Portable Ultrasound Scanner | General information: Multipurpose combined console/digital platform for general, abdominal, cardiological, musculoskeletal, vascular (including transcranial), pediatric (including neonatal) studies. Application: adult, pediatric and neonatal patients General specifications: Connection of at least 3 probes at the same time plus the probe for the Doppler CW. Color Monitor LCD/TFT of at least 17’’. Illuminated alphanumeric keyboard. Frequency range of at least 1 – 15 MHz. Dynamic range of at least 180 dB. Penetration depth of at least 30 cm. Frozen image zoom of at least 10X. Dynamic real time zoom of at least 4X. All software in English (as mandatory) and Russian, Ukrainian language (if available). The hardware and software included in the offer will allow the following application: cardiac; cardiac quantification abdominal; (including the study of abdominal vessels in Doppler modes), muscular; (including examination of the vessels of the extremities in Doppler modes) and examination of small organs. obstetrical and gynecological; in obstetrics (including advanced perinatology), including gestational sac mean diameter, femur length, crown-rump length, biparietal diameter and abdominal circumference, enabling estimation of gestational age and foetal weight (preferable); Small parts and tissues examination of the mammary glands, examination of the thyroid gland (including examination of the vessels of the neck in Doppler modes), peripheral and deep vascular; Research in angiology (including abdominal, peripheral and superficial vessels) Images Modes, at least: 2-D imaging M-mode B/M mode Dual 2-D/colour image mode cine loop CFI Doppler PDI (Power Doppler Imaging), PW (Pulsed Wave Doppler) CW (Continuous Wave Doppler) Needle enchancement ability Allowed measurements: body marks; length; angle; area; circumference; cavities; speed; flow. Software included functions, at least: trapezoidal images; vascular examinations; fetal weight indexes table; production and storage of images and videos; steering pulsed Doppler. Allowed probes: linear; convex; endocavity; transesophageal; microconvex; phased array; pencil. Communication and storage DICOM 3.0 conformity. Capability for database of patient images and information. Hard disk of at least 1 TB. At least one USB port and high-definition multimedia interface (HDMI) preferable. Probes to be included: Convex with at least triple frequency, bandwidth of at least 3Mhz, including 3.5Mhz frequency; (Used for Abdominal, OB/GYN, Abdominal Vessels, etc.), Supports Biopsy Adapters Linear with at least triple frequency, bandwidth of at least 2Mhz, including 6.5Mhz frequency; used for vascular, superficial and small organs, thyroid and breast, musculoskeletal), supports biopsy adapters Phased array with at least triple frequency, bandwidth of at least 2 Mhz including 3 Mhz frequency; Transvaginal transducer: 4.0-12 MHz (used for endovaginal, endorectal examinations) Accessories to be included: B/W printer for medical ultrasound applications. 5 liters of ultrasound gel. Compatible printing paper Image Enhancement Software Package: • (tissue or 2nd harmonic) for all sensors, • multi-angle composite scanning, to eliminate artifacts and shadows from hypoechoic structures, • Speckle-grain reduction algorithm, • contrast enhancement technology, • multi-beam scanning technology, • technology of automatic amplification of the color range Electrical Requirements: Voltage: 230V, 50 Hz Plug type: F Protections against overvoltage and overcurrent line conditions. All electrical connection according with the destination Country standards (Ukraine) Working time in battery mode and standard operations not less than 2 hours. Battery recharging time not more than 4 hours. Documentation requirements • Software licenses and user and maintenance manuals. Instructions for use and service manuals to be provided (including procedures for decontamination) • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local service agent. • Certificate of calibration and inspection to be provided • List to be provided of equipment and procedures required for local calibration and routine maintenance. • List to be provided of common spares and accessories, with part numbers. Other requirements. All standard accessories, consumables and parts required to operate the equipment, including all standard tools, cleaning and lubrication materials. At least two (2) years of full onsite warranty including 2 yearly preventive maintenances, full corrective maintenance including works and spare parts. Training on use, cleaning, disinfecting and basic maintenance on-site or in teleconference for at least 4 hours. Regulatory approvals: Compliance to - National Regulatory Agency/Authority (NRA) requirements compliance Free Sales Certificate And at least one of: - FDA 510k clearance (US FDA) - CE mark (EU), with indication of Notifying Body - Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan. Manufacturer authorization QMS ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes). General quality management (ISO 9001:2015 Quality management systems – Requirements). Application of risk management to medical devices (ISO 14971:2019 Medical devices – Application of risk management to medical devices). Safety and product standards: Declaration of Conformity by the Manufacturer is acceptable · ISO 29821:2018 Condition monitoring and diagnostics of machines – Ultrasound – General guidelines, procedures and validation. · IEC 60601-1:2005- Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. · IEC 60601-1-2:2014 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. · IEC 60601-2-37:2007 Medical electrical equipment – Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment. · IEC 61157:2007 – Standard means for the reporting of the acoustic output of medical diagnostic ultrasonic equipment. · IEC 60601-1-4 Medical electrical equipment – Part 1-4: General requirements for safety – Collateral standard: Programmable electrical medical systems. | 7 pcs |

Partial bids are allowed under this RFQ. The bidder shall not be required to quote for all Lots. However, Bidders are encouraged to quote for as many Lots as possible.

The goods are to be delivered maximum in 21 calendar days upon issuing of PO. The quotation shall be valid at least for 3 month after the closing date.

Email address of Contact Person in case of technical questions: seredenko@unfpa.org - Andrii Seredenko, SRH Project Manager.

The deadline for posting technical questions is Tuesday, July 4, 2023, 12:00 Kyiv time. Answers to questions in the form of a letter (by email) will be provided as soon as possible after this deadline.

Documents to be submitted with the bid:

- Completed and signed Quotation Form (financial bid);

- Technical bid, including brochures, pictures and product catalogue to demonstrate that specification and quality of the products are in line with the requirements listed in the bidding documents;

- FTP Questionnaire for Medical Devices;

- International and Ukrainian Certificates and documents (in addition if not specified in Product Description above):

- EC certificate (referencing the name/number of the notifying body) with an additional copy of EC Design Examination certificate, and/or 510k/PMA FDA clearance, and/or approval letter or certificate from national Ukrainian regulatory body (which confirms local registration of the product and permission to use in medical practice);

- A signed and dated document according to ISO 17050 stating compliance to critical ISO standards (e.g. ISO 13485 QMS) and directives, and which has a reference to the offered product. Proof of compliance to ISO standards in a form of copies of certificates shall be submitted.

- Letter of Attorney or Authorization Letter from the Manufacturer to the Bidder which authorizes sales and services, local registration, staff training in Ukraine.

If you are interested in submitting a quotation for these items, kindly fill in the attached Quotation Form and send it together with other required documents by email to the address indicated below:

| Contact person | Iryna Bohun |
| --- | --- |
| Email address | ua-procurement@unfpa.org |

Please submit your quotation in USD currency. Conversion of currency into the UNFPA preferred currency, if the offer is quoted differently from what is required, shall be based only on UN Operational Exchange Rate prevailing at the time of competition deadline.

The UNFPA terms of payment are 30 days after receipt of goods with shipping documents, invoices and other documentation required by the contract.

Your earliest response to this query would be highly appreciated, but not later than Monday, 10/07/2023, 13:00 Kyiv time.

Note: Current UNFPA supplier policies apply to this solicitation and can be found at: http://www.unfpa.org/suppliers.

Best regards,

Maksym Grechishnikov

Procurement and Logistics Associate

Email: grechishnikov@unfpa.org

Quotation Form

Name of Bidder:

Date of Bid:

Request for Quotation No:					UNFPA/UKR/RFQ/23/21

Currency of Bid price:						USD

Delivery time (weeks from receipt of order till dispatch):						
(Note: maximum number of days is: 21 days)

Expiration of Validity of Quotation (The quotation shall be 
valid for a period of at least 3 months after the Closing date.):	   3 months

Price Schedule:

| Lot and Item No. | Quantity, pcs | Unit price, USD, CPT (Kyiv) | Total price, USD, CPT (Kyiv) | Delivery schedule (weeks) |
| --- | --- | --- | --- | --- |
| 1. Colposcope | 7 |  |  |  |
| 2. 12 channel ECG device | 7 |  |  |  |
| 3. Portable Ultrasound Scanner | 7 |  |  |  |

In your offer, please include:

- Specific technical specifications of products offered;

- Quality standards of the products;

- FTP Questionnaire for Medical Devices

I hereby certify that this company, which I am duly authorized to sign for, accepts the terms and conditions of UNFPA (http://www.unfpa.org/resources/unfpa-general-conditions-contract ) and we will abide by this quotation until it expires.

Name and title								Date and Place

Fast Track Procurement Questionnaire 
for Medical Devices

PART I.  Manufacturer information

Bidder (if not manufacturer):				 Click here to enter text.

Manufacturer:   	Name of manufacturer:		 Click here to enter text.
			Country:			 Click here to enter text.
			Address (office):		 Click here to enter text.
			Address (manufacturing site(s)):	 Click here to enter text.
			Contact person’s name:		 Click here to enter text.
			Email:				 Click here to enter text.
			Phone:				 Click here to enter text.

PART II.  Product information

Product Identification (Trade name, Type, Model, Package size, Intended use, etc.):
 Click here to enter text.

Product Code, Reference number(s):  Click here to enter text.

Product details (materials, dimensions, size, volume, features, etc. For electrical devices specify voltage, frequency and plug supplied.): (E.g. If a stainless steel product, identify AISI type or composition. If a plastic product, identify type or composition.)
 Click here to enter text.

## PART III.  Regulatory Status

| 3.1 Is the product CE marked? Certification body and number: Click here to enter text. | □ Yes | Start Date: Click here to enter text. Expiry Date: Click here to enter text. |
| --- | --- | --- |
|  | □ No |
| 3.2 Is the product FDA approved? 510k clearance #: Click here to enter text. PMA clearance #: Click here to enter text. | □ Yes | Start Date: Click here to enter text. Expiry Date: Click here to enter text. |
|  | □ No |
| 3.3 Is the product approved by National Regulatory Agency or Department? Name of agency and type of approval: Click here to enter text. | □ Yes | Start Date: Click here to enter text. Expiry Date: Click here to enter text. |
|  | □ No |
| 3.4 Provide details of any other current regulatory approvals for this product. Name of jurisdiction and type of approval: Click here to enter text. | □ Yes | Start Date: Click here to enter text. Expiry Date: Click here to enter text. |
|  | □ No |

3.5         Manufacturer QMS ISO 13485 		Yes ☐ 	   No ☐
                                         QMS ISO 9001 		Yes ☐ 	   No ☐

  - Certification body and number: Click here to enter text.

  - Expiration date: Click here to enter text.

3.6  FOR STERILE PRODUCTS - If the manufacturing process is subcontracted:

| Name and address of the subcontractor | QMS certification of the subcontractor - Identify Regulatory body and/or number and expiry date |
| --- | --- |
| Click here to enter text. | Click here to enter text. |

3.7  FOR ELECTRICAL or BATTERY-OPERATED PRODUCTS

| If the device contains Lithium metal and Lithium ion batteries, does it comply with clause 38.3 of the recommendations on “Transport Of Dangerous Goods” from the United Nations? | Yes ☐ No ☐ |
| --- | --- |
| Does it comply with the latest IATA Dangerous Goods Regulations (DGR)? | Yes ☐ No ☐ |
| Testing laboratory, Test Report reference, specify standard | Click here to enter text. |

PART IV.  Checklist of required documentation

| Product class (EC MEDDEV) | Minimum documentation required Documents to be submitted must be true and valid copies. All documents submitted must be in English or be accompanied with certified translation. |
| --- | --- |
| class I (non-measuring, non-sterile and/or non-reusable surgical instrument, rsi) | ☐ Copy of ISO 13485* (or ISO 9001*) QMS certificate. ☐ A signed and dated Declaration of Conformity (DoC) according to ISO 17050 stating compliance to the relevant ISO standards and directives (for manufacturer), and which has reference to the offered product. ☐ Photo of the product and packaging (at various angles if necessary, preferably in a format where the dimensions and features can be visually verified from the photos). |
| class I measuring class I sterile class I rsi class IIa | ☐ Copy of EC certificate (referencing the name/number of the notifying body), and/or 510k FDA clearance, and/or approval letter or certificate from a National Regulatory Body. ☐ A signed and dated DoC according to ISO 17050 stating compliance to critical ISO standards (e.g. for sterilization, ISO 13485 QMS) and directives, and which has reference to the offered product. Note: If a sterilization activity is subcontracted to a third party, ISO 13485 QMS compliance is also required from the subcontracting company. ☐ Photo of the product and packaging (at various angles if necessary, preferably in a format where the dimensions and features can be visually verified from the photos). |
| class IIb class III | ☐ Copy of EC certificate (referencing the name/number of the notifying body) with an additional copy EC Design Examination certificate, and/or 510k/PMA FDA clearance, and/or approval letter or certificate from a National Regulatory Body. ☐ A signed and dated DoC according to ISO 17050 stating compliance to critical ISO standards (e.g. ISO 13485 QMS) and directives, and which has reference to the offered product. Proof of compliance to ISO standards in a form of copies of certificates shall be submitted if available. ☐ Photo of the product and packaging (at various angles if necessary, preferably in a format where the dimensions and features can be visually verified from the photos). |

*) UNFPA accepts the versions of currently active standards, which are recognized by the International Organization for Standardization at the time of document submission.

# Text boxes

Vendor’s Comments: